# Pouch sealing and seal inspection, sterile barrier packaging

EXAMPLE for a fictional company. Every spec value is an example value, not a recommendation. Replace it with the value from your drawing, specification or qualified procedure.

From coplain.com/industries/medical-device-contract-manufacturing: how to adapt it and common mistakes are on that page.

## Document control

| Document | WI-PK-014, Example |
| --- | --- |
| Revision | E, implements owner specification PS-208 Rev C |
| Scope | Sealing pre-made pouches on sealer SL-02 in the controlled packaging area |
| Validated parameters | Temperature, dwell time and pressure per validation report PV-SL02-03 |
| Approved by | Quality engineer, production manager; owner approval on file |

## Tools, materials and PPE

- Pouch sealer SL-02 with current calibration label

- Lot traveler and device history record sheet

- Seal inspection light box

- Gloves and garments per the area gowning instruction

## Steps

| No. | Step | Key point | Reason |
| --- | --- | --- | --- |
| 1 | Check the sealer settings. Compare the displayed temperature, dwell and pressure with the validated values in the header and record them on the DHR sheet. | If any value is outside its validated range, do not seal; call the quality engineer. | Seals made outside the validated window are not covered by the validation. |
| 2 | Run the start-up seal. Seal one empty pouch and inspect the seal on the light box. | Continuous seal width, no channels, wrinkles or burns. SAFETY: The sealing jaws are hot; keep fingers clear of the jaw area. | Confirms the sealer is working before product is packed. |
| 3 | Seal product pouches. Place the device in the pouch with the label side up and seal within the marked zone. | Product must not sit in the seal area. | Product in the seal can create a channel that breaks the sterile barrier. |
| 4 | Inspect every seal. Inspect each seal on the light box and set aside any pouch with a defect. | Rejected pouches go to the red bin with a reject tag, never back into the line. | Keeps nonconforming product identified and segregated. |
| 5 | Complete the DHR. Record lot number, quantity sealed, quantity rejected and your initials on the DHR sheet. | Write entries at the time of work, in ink. | The device history record proves how this lot was made. |

## Acceptance criteria

- Settings recorded and within validated ranges

- Every seal continuous with no channels, wrinkles or burns

- DHR complete with lot, quantities and initials

## Sign-off

| Prepared by |  | Date |  |
| --- | --- | --- | --- |
| Approved by |  | Date |  |
| Trained / signed by operator |  | Date |  |